## Slide 1: Medical Ultrasound: Fundamental Physics & Safety Implications

## Slide 2: Basic Principle of Ultrasound
Waves of sound are sent into the body and ‘bounce’ off structures
The returning echoes are then collected & used to produce an image

## Slide 3: Basic Principle of Ultrasound
Sound waves are produced within the transducer (Piezoelectric effect)
Sound propagates into tissue when the transducer is applied to the patient. Coupling gel is required to combat air between interfaces.
Returning sound is captured by the transducer, changed to electrical signal and passed to the processing unit for display on the monitor.

## Slide 4: Piezoelectric Effect
An electric current is briefly passed over the layer of crystals present within the transducer
The current physically alters the shape of the crystal
Once the current is switched off , the crystal snaps back to it’s normal shape producing a sound wave
This cycle is repeated to produce continuous scanning

## Slide 5: Piezoelectric Effect
Electrical charge applied – compression of crystals
Electrical charge off – rarefaction of crystals
Normal crystal
Sound wave emitted

## Slide 6: Acoustic impedance (z)
Refers to the resistance offered by the tissue to the travelling sound wave
Determines how much sound is transmitted forward and how much is reflected back
Determined by the density of the tissue (p) and the speed of sound (c) within that tissue type
Z= pc
Interface between tissue
The amount of sound allowed to pass through a border of two differing tissue types will depend on how closely matched the impedance values are – i.e. the greater the mismatch the less sound throughput
(reflected sound)
(transmitted sound)
Original sound wave

## Slide 7: Sound propagation
The more alike the neighbouring tissues are the more sound is transmitted
This is known as the intensity reflection co-efficient (R)
Determines the ratio of energy reflected – where 0 = complete transmission and 1 = complete reflection
Tissue borders | Intensity reflection coefficient (R) | Sound reflected - percentage
Soft tissue/ water | 0.002 | 0.2
Fat/muscle | 0.0108 | 1.08
Bone/muscle | 0.412 | 41.2
Bone/fat | 0.49 | 49.0
Air/soft tissue | 0.999 | 99.9

## Slide 8: Reflection
When a large amount of sound is reflected this produces a brighter echo and appears hyper echoic on screen
Trade off -there is very little imaging capability past this strong reflector
  E.g. acoustic shadow posterior to bone

## Slide 9: Reflection
The strength of the sound reflected back to the transducer will also depend on the type of interface it meets
A smooth surface will return a stronger echo (specular reflection)
An irregular surface will cause the beam to scatter into more, smaller echoes -therefore reducing the overall strength (diffuse reflection)

## Slide 10: Attenuation
The process by which a sound beam will lose it’s intensity
Reflection is one of the main methods of attenuation
A beam is attenuated on it’s way in to the body but also on it’s return path back to the transducer
Directly proportional to frequency – i.e. higher frequencies attenuate more therefore cannot travel as far. Not generally a problem for superficial MSK scanning.
  Reflection
    Forming of echo
  Absorption
    Transformation of sound energy into heat energy – stored within tissue
  Scattering
    Form of reflection where the majority of the echoes are not returned to the transducer –spread throughout the organ

## Slide 11: Attenuation
  Refraction
    Change of beam pathway – occurs at tissue borders where there is a mismatch in speed of sound causing a change in wavelength
  Divergence
    The widening of the beam from origin – results in a loss of intensity
Wider beam at the far field than the near field
Tissue 1
Tissue 2
Snell’s law – formula that expresses the angles of incidence (i) and refraction (r)
sinθi = c1
Sinθr c2
i
r

## Slide 12: Image generation
Returning echoes are received by the piezoelectric crystal within the transducer
Crystal again is vibrated– this time converts the vibration to an electrical signal
The processing unit within the machine recognises the position of each crystal – i.e. this will allow the computer to plot the echo on the screen depending on where on the transducer it was received and the strength of returning signal

## Slide 13: Image generation
The area of interest lies within the left field. As this is detected by crystals towards the left of the transducer this has been correctly plotted on the screen

## Slide 14: Image generation
The system must additionally plot the correct depth an echo has returned from
The time taken for a returning echo to be detected will determine how far it has travelled
  E.g. a longer time will translate to a greater depth displayed on screen

## Slide 15: Image generation
Time for pulse and receive X seconds
Time for pulse and receive 2X seconds

## Slide 16: Image generation
The system plots every echo this way
Bases it’s calculations on several assumptions
  Path of an echo is straight – original or returning
  The speed of sound (1540 m/s)and attenuation is constant through all tissue
  All echoes have originated from the centre of the beam
  Returning echoes – time taken for these to be detected is directly proportional to the distance travelled

## Slide 17: Image generation
These assumptions can lead to false appearances being displayed on the image
These are known as artefacts
Common artefacts include
  Acoustic shadow
  Acoustic enhancement
  Reverberation
  Edge shadow
  Mirror image

## Slide 18: Image display
Each echo is displayed as a grey scale
Level of grey determined by the amplitude of the echo received
  Strong echo = white
  No echo = black
Approximately 64 grey scales available on most scanners – can alter manually

## Slide 19: Image display
The ability of the ultrasound machine to distinguish different structures and display these correctly is known as resolution
Several types are important in MSK Ultrasound imaging:
  Spatial
    Axial
    Lateral
Spatial resolution – the ability to distinguish two structures which lie very close together
Axial
Lateral

## Slide 20: Resolution
Axial resolution
  Determined by pulse length – i.e. short pulse length = greater axial resolution
  Affected by the frequency – higher frequencies have shorter wavelengths and pulse lengths
Lateral resolution
  Determined by the beam width – i.e. beam must be narrow enough to ‘see between’ two structures and register them separately
  Affected by the use of focus

## Slide 21: Resolution
Axial
Lateral

## Slide 22: Resolution
Temporal resolution
  Ability of the machine to correctly distinguish two separate actions occurring at separate times and display them accordingly – i.e. the number of times a section of tissue is scanned and displayed
  Faster the update (frame rate) = higher temporal resolution – REAL TIME
Sector width – the narrower the section the faster the whole area can be imaged = more times (increased frame rate) and better temporal resolution
Focal zones – more focal zones require longer sweeps and thus have lower temporal resolutions
3 zones – requires 3 sweeps
1 zone – 1 sweep

## Slide 23: Knobology
Typical machine – important controls to increase image quality/ detail
TGC – used to alter amplitude of echoes at different levels of the image – e.g. gradual slope useful to even out echoes from deep structures with those superficially. Differs from overall gain which amplifies all echoes equally
Also known as contrast – alters the number of grey scales available for the computer to use for display
Sector width – alters the number of scan lines in an image
Frequency – alters the resonant frequency employed – higher for more superficial structures = better detail
Depth – distance able to be imaged

## Slide 24: Doppler imaging
Colour – demonstrates the presence and direction of flow. Dependent on incident angle of beam.
Spectral – allows calculations of velocity and pulsatility to be calculated. Dependent on incident angle of beam.
Power – demonstrates the presence of flow; gives no indication of direction. More sensitive for low flow. Not angle dependent and used widely in MSK field.
colour
spectral
power

## Slide 25: Doppler imaging
Frequency of a sound wave is changed when it encounters a moving object
  Increases when object moves closer to the wave
  Decreases when object moves further from the wave
Proportional to the velocity of the moving target – Doppler shift
Doppler shift equation

## Slide 26: ALARA
As Low As Reasonably Achievable
Means use the least amount of ultrasound to gain a diagnostic image
Overarching principle of imaging – in ultrasound this translates to
  Lowest output power
  Shortest scanning time possible
Use of supplementary imaging techniques (e.g. Doppler) increases the output power.

## Slide 27: Bio-effects of ultrasound
No evidenced adverse effects of ultrasound in human subjects
Theoretical effects fall into two categories
  Mechanical
    Stable cavitation
    Non-stable cavitation
  Thermal
Cavitation –
  the creation of an empty space within tissue

## Slide 28: Mechanical bioeffects
Stable cavitation
Tiny gas bubbles in tissue
Pressure of the sound wave will cause a transient change in their size
Considered relatively safe
Compression
Rarefaction
Gas bubble shrinks during peak positive pressure
Gas bubble swells during peak negative pressure

## Slide 29: Mechanical bioeffects
Unstable cavitation
When there is high enough intensity within the wave to cause the gas bubbles to burst and collapse
Will cause damage to surrounding material – i.e. causes an ‘empty space’ to appear where the gas bubble once occupied
Proven cases in rodent lung
High intensity beam
Average intensity beam
Normal oscillation – stable cavitation
Gas bubble bursts and destroys the surrounding tissue – creates a hole

## Slide 30: Thermal bioeffects
The heating of tissue through which a sound wave passes – process of absorption
Highest risk of heating within tissue with high absorption coefficient e.g. bone
Also affected by higher frequencies and screening times

## Slide 31: Safety
Ultrasound operators are responsible for minimising the risk of bioeffects
Manufacturers are required to display information on the live image that will help operators make informed decisions – safety indices
  Thermal Index (TI)
  Mechanical Index (MI)

## Slide 32: TI
Determines the ratio of the power produced to the power estimated to raise the temperature of surrounding tissue by 1°C
TI = W
W deg
i.e. TI = overall power
power needed to raise
temp by 1°C

## Slide 33: MI
Estimates the potential for cavitation by the formula below
MI = PNP
√f
i.e. the peak negative pressure divided by the square root of operating frequency

## Slide 34: Safe limits
MI Level | Concern
MI > 0.3 | Possible risk of slight damage to the neonatal intestine/lung tissue
MI > 0.7 | Theoretical risk of cavitation. Risk of cavitation in conjunction with contrast agent
TI Level | Concern
TI > 1 | Risk of increased tissue temperature
TI > 3 | Significant risk of increased tissue temperature
The British Medical Ultrasound Society (BMUS) have previously released recommendations on MI and TI values
(Adapted from Gibbs et al, 2010)

## Slide 35: Controlling acoustic output
The TI and MI values can be controlled by
  Non- Stationary Probe – if specific tissue is continually interrogated this can cause localised heating.
  Frequency – higher frequency will cause increase in TI due to higher absorption
  Focus – As the intensity is highest at the focus, the site and number of focus points along a beam will affect the MI/TI values

## Slide 36: Indicative reading
British Medical Ultrasound Society (2009) Guidelines for the safe use of diagnostic ultrasound equipment https://www.bmus.org/static/uploads/resources/BMUS-Safety-Guidelines-2009-revision-FINAL-Nov-2009.pdf
Gibbs, V., Cole, D. and Sassano, A. (2009). Ultrasound physics and technology. Edinburgh: Churchill Livingstone/Elsevier.
Kremkau, F., Forsberg, F. and Kremkau, F. (2011). Sonography principles and instruments. St. Louis, Mo.: Elsevier/Saunders.